## Slide 1: Eclipse Algorithm Notes

## Slide 2: CT data
The conversion is done following CT calibration curves configured in Beam Configuration. Multiple CT scanners can be configured in the system and each CT scanner has its own set of calibration curves.
For each CT scanner the following curves are available:
■ Electron Density: This curve is used by the AAA and MRDC algorithms to convert HU values to relative electron density.
■ Mass Density: This curve is used by the Acuros XB and eMC algorithms to convert HU values to mass density. From the derived mass density, Acuros XB and eMC determine the material composition of voxels in the image.

## Slide 3: CT Data Scanner
For each CT image imported into the system, the CT scanner name in the image header is compared with CT scanners stored in the database.
If a match is found, the custom CT calibration curve is used for all dose calculations.
If a matching CT scanner is not found, you need to manually define the CT scanner to use for the image series

## Slide 4: AAA handles CT data
The AAA algorithm truncates all HU values higher than the maximum HU value defined in the CT calibration curve to the maximum electron density value. A warning is shown in the DCF status window and logged in the field calculation notes when such a situation occurs.
Additionally, the maximum electron density modeled in AAA is 15. Electron densities larger than that are truncated to the maximum value (that is, the value 15 is used).

## Slide 5: Accuracy
Gamma error is defined as the shortest distance in four dimensional space where the relative dose represents the fourth dimension. For gamma metric, the unit of the distance and the dose are rescaled: dose is in units of a percent of the maximum dose, and the distance in units of given number of mm.

## Slide 6: Accuracy (cont.)
Dose calculation accuracy goal defined by the gamma error scale is considered to be fulfilled when gamma error value for 95% of the examined measurement points is no more than 1.

## Slide 7: MU Calcs
Based on output factor measurements performed for different field sizes in a certain reference geometry and calibration calculations made for the reference field size.
The change in the output factor as a function of field size is caused by changes in phantom scatter, head scatter and collimator back scatter into the monitor chamber.

## Slide 8: MU Calcs (cont.)
The phantom and head scatter effects are accounted for by the photon beam source model and the volumetric dose calculation algorithm (AAA or Acuros XB). The remaining change in the output factors is assumed to be caused by collimator back scatter.

## Slide 9: MU Calcs (cont.)
The final MU are calculated from the prescribed dose, plan normalization, field weight, field normalization and a normalization factor determined by the dose calculation algorithm.
The normalization factor determined by the AAA and Acuros XB is the MU value for 1 Gy to 100% of the current field. AAA and Acuros XB calculate the monitor units at the normalization point MUnorm

## Slide 10: Normalization

## Slide 11: Calculation grid
AAA: 1-5 mm dose calc grid
Acuros: 1-3 mm
AAA calculates the dose first into a divergent dose matrix.
The width of the divergent dose matrix depends on the jaw positions.
An additional margin is added in order to calculate also the low dose tail resulting from head scatter and phantom scatter.
  This margin depends on the field size and the selected calculation resolution. The default margin size is 12 cm. However, the margin is reduced so that the number of calculation points at isocenter plane does not exceed 74000
Margin is always at least 7 cm; so between 7-12 cm; for AAA dose attains 0 values when outside the margin

## Slide 12: Calculation grid
In the axis perpendicular to the image slices, AAA and Acuros XB adapt the grid resolution to ensure that the dose is always calculated exactly on the image slices.
If the slice thickness is larger than the defined grid size, AAA and Acuros XB may calculate the dose on dose planes between the image slices. If the slice spacing is smaller than the defined grid size, AAA and Acuros XB may skip calculating the dose on some slices.
For image slice thickness IS, AAA and Acuros XB calculate IS/ n...IS/3, IS/2, IS, IS × 2, ΙS × 3...ΙS × n and sets the dose plane spacing to the value closest to the defined grid size

## Slide 13: Angular Resolution in Conformal Arc and VMAT Calculations
This option specifies the angular resolution in degrees for arc field calculations.
It integrates the existing machine control points into beams placed along the arc at the defined interval. For conformal arc and VMAT fields, OFF means that the beams used to calculate the arc coincide to the machine control points.
For standard arc fields, OFF means that the angular resolution is 5 degrees.

## Slide 14: Progressive Resolution
The initial conditions for the PO algorithm are defined using control points to represent each VMAT field.
The algorithm uses multi-resolution approach to optimize the plan. This approach was first described in the research on volumetric modulated arc therapy. This means that the dose is modeled using first a lower number of dose calculation segments that are distributed evenly in each field.
The number of dose calculation segments increases when moving from one multi-resolution level to another.
The dose in a dose calculation segment is calculated from the combined fluence through the MLC apertures at the control points located within a certain sector of the arc.
Leaf motion is modeled by interpolating leaf positions between the control points.
Leaf tongues are modeled by modifying the MLC aperture outline to effectively account for the tongue-and-groove effect.

## Slide 15: Progressive Resolution
The angle resolution of the dose calculation segments gets more accurate as the optimization progresses, and in consequence, the dose also gets more accurate.
The number of control points remains the same during the whole optimization.
At the beginning of the optimization, the initial MLC shapes are conformed to the targets and the initial dose rates are equal for all dose calculation segments.
The MLC shapes and dose rates of the different control points in the VMAT field are optimized.
During the initial phases of the optimization bigger adjustments are made in leaf sequencing. The size of these adjustments decreases as the optimization progresses through the levels.
During the optimization, the algorithm proceeds through multi-resolution levels progressively increasing the accuracy of the dose calculation.
At the first multi-resolution level, only a few dose calculation segments are used to model the dose, and each multi-resolution level contains progressively more dose calculation segments.
The angle between the resulting dose calculation segments on the last multi-resolution level (4) will be approximately 2°-4°.The total number of dose calculation segments used depends on the span of the arc.

## Slide 16: Primary Source Model (photon beam)
Point source located at target plane
Physical effects of finite size modeled by the effective target spot size parameters
Models brem photons created in target that do not interact w/ treatment head
Initial photon spectra after target on beam central axis were simulated using BEAMnrc using realistic target materials and thicknesses
Beam hardening in flattening filter modeled by attenuating the initial spectrum by radially varying amount of the flattening filter material.
Separate energy spectra for every fanline of the broad beam are derived from the mean energy curve. Non-uniform photon energy fluence is modeled by radially varying intensity profile curve

## Slide 17: Second Source
Gaussian plane source located at bottom of flattening filter
Models photons that result from interactions in the accelerator outside the target (primarily in flattening filter, primary collimators, and secondary jaws)
Due to lower location of second source, radiation from second source diverges more than radiation from primary source (effect most noticeable outside the beam defined by the primary source)
Disabled for FFF beams; also disabled for hard wedges

## Slide 18: Electron Contamination
Dose deposited in build-up region not accounted for by primary and extra-focal photon components
Also used to model photon contamination (photons created in e- interactions)
Modeled using radiological depth dependent curve that describes total amount of electron contamination dose at a certain radiological depth
The electron contamination is modeled with a depth-dependent curve that describes laterally integrated electron contamination dose at different depths
The shape of the electron fluence is obtained as a convolution of the aperture shape and a 2D kernel consisting of two Gaussian components

## Slide 19: Phase Space Model
Physical jaws are assumed to have zero transmission
The phase space consists of a primary source phase space and one or more non-primary source phase spaces.
The primary phase space consists of a two-dimensional energy fluence grid and a two-dimensional grid of spectra.
The non-primary phase space consists of a three-dimensional energy fluence grid (that is, a two-dimensional energy fluence grid that varies with distance from source) and a single spectrum. For an arc field, the source model calculates several static beam phase spaces.

## Slide 20: Modeling Primary Source
Inputs:
  Primary opening ratio matrix (static jaw/MLC opening or actual fluence for an IMRT field / arc segment)
  Photon energy spectrum
  Mean Energy
  Intensity Profile

## Slide 21: Second Source Parameters
The second source model includes a spectrum whose energy axis is scaled to obtain a given mean energy.
Off-axis variation in the second source spectrum is not modeled.
The relative intensity of the second source compared to the primary photon source, the width of the Gaussian at the source plane (size of the second source), and the mean energy are free parameters.

## Slide 22: Mean Energy
Another important parameter that affects the energy spectrum used by the source model is the mean energy as a function of the radius from the beam CAX.
This curve is used by the source model to determine the beam hardening effect of the flattening filter on the photon spectrum. Based on the mean energy curve and the user-specified flattening filter material, the source model determines the energy spectrum of the beam as a function of radial distance from the beam CAX.
For flattening filter free (FFF) beams, the mean radial energy curve models the off-axis variation in the initial bremsstrahlung spectrum.

## Slide 23: Beam Modifiers in the Treatment Beam Model
Most beam modifying accessories affect only the beam fluence used in dose calculation.
Blocks, trays and MLCs employ a user-defined transmission factor to model the radiation transmitted through the accessory.
Compensators, Dynamic Wedges and IMRT fields also modify the fluence of the beam.
The head scatter effects are taken into account by using a finite size second source located at the bottom surface of the flattening filter

## Slide 24: AAA Accuracy in Lung
For 4 MV to 6 MV energies and field sizes larger than or equal to 5×5 cm2, AAA tends to underestimate the dose in lung and overestimate the dose in water equivalent tissue after the lung.
  For 6 MV, the errors are smaller than 3% of the field CAX dose maximum. However, error from local dose can become significantly larger, if the local dose is small relative to the dose maximum.
For 10 MV to 20 MV energy modes and field sizes smaller than or equal to 5×5 cm2, AAA tends to overestimate the dose in lung. The errors become larger as the field size decreases and the lung density decreases.
  For 18 MV and 3×3 cm2 field, the maximum error in lung with density 1.85 g/cm3 is approximately 7% of the field dose maximum.

## Slide 25: Accuracy of AAA and Acuros XB for static MLC fields
For 18 MV, AAA and Acuros XB tend to underestimate the dose at shallow depths for certain static MLC shapes.
  At 3 cm depth, the error can reach 3% of the local dose.
For 6 MV, AAA and Acuros XB may underestimate the dose at large depths (larger than or equal to 20 cm) for certain static MLC shapes.
  The errors are more pronounced for elongated (for example H-shaped) openings. The errors can reach up to ~4% of the local dose

## Slide 26: AAA
3D pencil beam convolution/superposition algorithm
Uses separate Monte Carlo derived modeling for primary photons, scattered extra-focal photons, and electrons scattered from the beam limiting devices.
The lateral dose deposition characteristics are modeled with six exponential curves.
The functional shapes of the fundamental physical expressions in the AAA enable analytical convolution, which significantly reduces the computational time.

## Slide 27: AAA
The AAA accounts for tissue heterogeneity anisotropically in the entire three dimensional neighborhood of an interaction site, by using photon scatter kernels in multiple lateral directions (Scatter Kernels).
The final dose distribution is obtained by the superposition of the dose calculated with photon and electron convolutions.

## Slide 28: Scatter Kernels
The scatter kernels describe the phantom-scatter effects for different beam qualities.
The EGSnrc Monte Carlo code was used to compute scatter kernels for mono-energetic pencil beams in water.
A poly-energetic scatter kernel is constructed as a weighted sum of the mono-energetic scatter kernels.
During the 3D dose calculation these kernels are scaled according to the densities of the actual patient tissues determined from the CT images

## Slide 29: Beamlets
The dose calculation is based on the convolutions over the beamlet cross-sections separately for the primary photons, extra-focal photons (second source), scatter from hard wedges, and for electrons contaminating the primary beam
All depth-dependent functions used in the beamlet convolutions are computed along the central fanline of the beamlet using the depth coordinate z.
Lateral dose scattering due to photons and electrons is defined on the spherical shell perpendicular to the central fanline of the beamlet.
The AAA makes the assumption that the dose resulting from photon and electron scatter can be calculated by a division into two main directions, which are the lateral and the depth scatter.

## Slide 30: Beamlets

## Slide 31: Volumetric Dose Calculation
For volumetric dose distribution calculation, the patient body volume is divided into a matrix of 3D calculation voxels based on the selected calculation grid
The geometry of the calculation voxel grid is divergent, aligning the coordinate system with the beam fanlines.
Every calculation voxel is associated with the mean electron density ρ that is computed from the patient CT images according to a user-specified calibration curve.

## Slide 32: Photon Dose Calculation – Energy Based
The dose convolution in the AAA is performed in terms of energy.
In comparison with dose-based convolution, which is more common, energy convolution allows the energy to be conserved more accurately even in more complex heterogeneous convolutions.
The energy is converted to dose using the scaled-water approximation.

## Slide 33: Lateral Density Scaling of the Photon Scatter Kernels
Density scaling of the individual pencil beams is done by scaling the energy at each location by the average density between the calculation point and the origin of the pencil beam.
In practice, this is done by dividing the kernel of the following equation into a finite number (16) of rays that emerge from the origin.

## Slide 34: Dose Optimization Algorithms
Photon Optimization algorithm (PO): Determines the optimal field shape and intensity by iteratively conforming the dose distribution to the desired objectives until an optimum solution is reached. Optimizes static field IMRT plans, VMAT (RapidArc) plans, Elekta VMAT plans and Siemens mARC plans based on optimization objectives.
Dose Volume Optimizer algorithm (DVO): Determines the optimal field shape and intensity by iteratively conforming the dose distribution to the desired objectives until an optimum solution is reached.

## Slide 35: Dose Optimization Algorithms

## Slide 36: Dose Calculation
The dose calculation from the fluences is based on multi-resolution 3D convolution of Monte-Carlo-generated point-spread function kernels
For the fluence-based static field IMRT optimization, the dose calculation is straightforward.
The VMAT optimization calculates fluence for dose calculation from a dose calculation segment, in which the MLC configuration and dose rates are converted into a fluence. The fluence models the leakage caused by the rounded leaf ends, tongue-and-groove effect and the transmission through the leafs

## Slide 37: Cost Function
If the dose-volume objectives are not met, a weighted quadratic cost is added to the total objective function.
For the upper objective, the cost is applied for the portion of doses that exceed the desired dose value and volume level.
For the lower objective, the cost is applied for the portion of doses that fall short of the desired dose value and volume level.

## Slide 38: Smoothing Objectives
The fluence needs to be smooth to enable leaf motion calculations with the LMC.
Fluence smoothness is ensured by adding an objective that includes the difference between the neighboring fluence values (X smooth and Y smooth). This objective is linear up to the value of 3% of the maximum fluence value, where it saturates to a constant value.
The saturation non-linearity allows for large fluctuations in the fluence where required.
Typically, it is more important to have a smoother fluence in the X-direction to ensure the minimal MU factor for the LMC.
Smoothing objective is available only for static field IMRT optimization with DVO and PO.

## Slide 39: MU Objective
The MU Objective can be used to control the number of MU that the PO or PRO optimizer produces.
Minimum and maximum values can be defined.
An extra multiplier is applied to the total objective function value if the number of MU is not in the desired range.
Strength value can be used to modify the strength of the effect.

## Slide 40: Normal Tissue Objective
The Normal Tissue Objective is used for the part of the body which does not include the PTV to limit the dose level and prevent hot spots in healthy tissue.
In addition, the Normal Tissue Objective can be used for obtaining a sharp dose gradient around the PTV.
Normalization of the Normal Tissue Objective is done as follows: level 1.0 (100%) corresponds to the lowest upper objective defined for the target. If no upper objective is defined, level 1.0 (100%) is 1.05 times the highest lower objective defined.

## Slide 41: Normal Tissue Objective Shape
The shape of the Normal Tissue Objective is controlled with the following parameters:
■ Distance from PTV border ( Xstart )
■ Start dose ( f 0 )
■ End dose ( f∞ )
■ Fall-off ( k )

## Slide 42: Handling of Large Fields in PO and PRO
The maximum field size in x direction that can be treated with a single carriage group is based on the maximum leaf span for the MLC, reduced by 0.5 cm due to the need to park closed leaves under the jaws.
  For example, if the maximum leaf span for an MLC is 15 cm, the maximum field size in x direction becomes 14.5 cm.
If the field size in x direction is larger than 14.5 cm, the field can contain unnecessary field openings.
In order for the PRO and PO to close these openings as much as possible within leaf speed limits, the underlying structure has to have dose objective (for example NTO) that minimizes extra radiation.
  If no objective is defined that prevents radiating a non-target volume, PRO and PO might not close all the field openings.

## Slide 43: Photon Optimization Algorithm (PO)
The PO combines the previous optimization methods used for static field IMRT and arc field IMRT with DVO and PRO.
The main difference of the PO algorithm from the earlier optimization algorithms (DVO and PRO) is that the earlier optimization algorithms used a point cloud model for defining structures.
  The PO algorithm uses a new structure model, where structures, DVH calculation and dose sampling are defined spatially by using one single matrix over the image

## Slide 44: Structure Model
The DVH for the structure is evaluated using volume weights defined for each voxel.
The volume weight of the voxel defines the ratio of the original structure segment inside the voxel.
For small structures, the DVH is super-sampled from the dose matrix to make the DVH look smoother.
In order to minimize the run time memory consumption, the matrix size is first set to the size of body bounding box.
Then the field field-of-views are formed, and the final matrix size is set to contain only the part of the body bounding box that can receive radiation from the fields.

## Slide 45: Photon Optimization Algorithm and IMRT
For IMRT, the PO determines the optimal field shape and intensity by iteratively conforming the dose distribution to the desired objectives until an optimum solution is reached.
For each field, the PO fits the fluence to the target projection with 5 mm margin. Then the created fluence object is expanded symmetrical to the field isocenter (by adding fluence pixels with 0 value). The maximum size of the fluence object is 40×40 cm.
The size of the fluence object determines the region where the user can edit the fluence.

## Slide 46: Photon Optimization Algorithm and IMRT (cont.)
The PO performs the optimization for an IMRT plan as a minimization problem using simple gradient optimization with line minimization.
Initially, all the fluences are zero, or, alternatively, the fluences from a previous optimization can be used as the initial guess.
The optimization modifies these fluences in each iteration and calculates the dose from the fluences after each modification
Once the doses at the points of the point clouds representing the patient volumes are evaluated, the objectives at the points and the derivatives of the point objectives can be calculated.
The cost functions are evaluated for each point in each volume.
The derivatives of the costs at each point are back-projected to the fluences, forming the gradient.

## Slide 47: Photon Optimization Algorithm and IMRT (cont.)
The PO uses the gradient search method. The gradient search is divided into two phases; gradient evaluation and line search.
  Gradient evaluation generates the gradient direction and the gradient length
  Line search evaluates the objectives on a line segment along the gradient and finds the minimum along the line segment.

## Slide 48: Intermediate Dose
The PO can use a calculated plan dose as an intermediate dose when optimizing an IMRT plan.
The algorithm calculates the difference between the intermediate dose and the first round optimization result and uses this difference to compensate the optimization result in the consequent iterations.
If a new intermediate dose is calculated after the first optimization iteration, the difference is calculated again and it will be used to compensate subsequent iterations.
Using an intermediate dose is particularly useful if the DVH calculated during optimization deviates from the DVH produced during dose calculation, for example when there is a lot of heterogeneity in the volume to be treated.
For VMAT, the second optimization round starts at the last multi-resolution level (level 4).

## Slide 49: Photon Optimization Algorithm and VMAT
The PO creates VMAT (RapidArc, Elekta VMAT, Siemens mARC) plans based on dose-volume objectives. VMAT fields use DMLC, variable dose rate and variable gantry speeds.
The PO algorithm generates a sequence of control points which define MLC leaf positions and MU/deg as a function of gantry angle. MU/deg is encoded in DICOM and the Varian system database with the cumulative meterset weight, which defines the increase in MU between control points relative to the total MU in the field.
This information is transferred to the treatment machine as such, and the machine control system determines how dose rate and gantry speed will be modulated to deliver the plan.
The PO algorithm uses an objective function to optimize the plan and to evaluate its quality. The objective function is the sum of the dose-volume and other user-defined objectives.

## Slide 50: MR Levels
Inside each multi-resolution level there are several steps. Each step has its own internal calculation parameter set.
The optimization allows some discontinuities in the delivery during early phases of the optimization, and decreases the size of the discontinuities stepwise as the optimization progresses.
At the step borders, the delivery is forced to be within specified discontinuity levels. This may increase the objective function values when moving from one step to the next, and a peak in the objective function curve may be seen.
The number of steps in different multi-resolution levels varies.
Due to the nature of the optimization process, the PO algorithm is not fully deterministic. Therefore successive optimizations with the same constraints may yield different results

## Slide 51: Avoidance Sectors
Ranges of gantry rotation with a zero delivered dose rate.
Can define up to two avoidance sectors for each arc field in the optimization.
The minimum length for an avoidance sector is 15 degrees.
  Similarly, the minimum length for the beam-on sector between two avoidance sectors, or between gantry start/stop angles and an avoidance sector is 15 degrees.

## Slide 52: Jaw Tracking for VMAT
Jaw tracking dynamically moves the collimator jaws during beam-on to keep them as close to the actual MLC aperture as possible.
  This reduces leakage between the MLC leaves.
The initial user-defined collimator jaw positions for the plan are used as the maximum limit for the jaws.
Jaw tracking does not move the collimator jaws outside this maximum limit.

## Slide 53: Collimation Mode
The collimator motion mode in RT Administration can have the following values:
■ Empty: Not allowed if jaw tracking is requested.
■ Static: Collimator does not move during the arc delivery.
■ Dynamic: Collimator moves within the set maximum speed limit during the arc delivery.
Motion modes must be the same for both jaws on the axis.
Collimator motion modes override the Field X setting earlier used for jaw tracking.

## Slide 54: Dose Volume Optimizer (DVO) Algorithm
Determines the optimal field shape and intensity by iteratively conforming the dose distribution to the desired objectives until an optimum solution is reached
For each field, the DVO fits the fluence to the target projection with 5 mm margin.

## Slide 55: General Features of Fluence Delivery Modeling Algorithms
In Eclipse, the final dose calculation for all MLC fields is based on fluences.
Features that are taken into account in all fluence delivery modeling algorithms include leaf transmission, dosimetric leaf gap, and tongue-and-groove modeling.
They are used in the final fluence calculation for all fields that contain an MLC, such as static MLC fields, IMRT fields and all arc fields

## Slide 56: Dosimetric Leaf Gap in Fluence Delivery Modeling Algorithms
The fluence delivery algorithms deal with the rounded leaf end transmission by means of the dosimetric leaf gap configuration parameter.
The algorithms model the shape of the leaf edges as sharp (as opposed to rounded) and take the rounded leaf transmission into account by shifting the leaf tip positions in the actual fluence calculation.
Leaf tips are shifted by pulling each of them back by half the value of the dosimetric leaf gap parameter so that the gap between a fully closed leaf pair equals the dosimetric leaf gap parameter

## Slide 57: Tongue-and-Groove Modeling in Fluence Delivery Modeling Algorithms
An exposed tongue in a field modifies the delivered fluence by blocking some of the additional radiation.
The amount of blocking is proportional to the ratio between the tongue and leaf widths. (The groove also modifies the fluence, but that effect is smaller and is not modeled in the algorithms.)
The tongue-and-groove effect is modeled in the fluence delivery algorithms by extending the leaf projections in the direction perpendicular to the leaf motion with an approximate extension parameter, which is slightly smaller than the real tongue width. This parameter is not user-configurable, and it depends on the MLC model and algorithm version.
The tongue-and-groove effect is taken into account in all fluences used in dose calculation and is visible in the actual fluence shown for IMRT fields.
The effect is more significant in DMLC and arc treatments than static MLC delivery techniques.

## Slide 58: Sliding Window Method
The main idea of the Sliding Window method is to sweep MLC leaf pairs of varying aperture sizes and speed over a field during the beam-on.
Implementations of the Sliding Window method for field widths under the Effective Leaf Out of Carriage Distance (ELOC) use the following strategy:
■ Start of delivery : Leaf pairs are at the left border of the target. Leaf pairs exposing the target to a normal or large amount of radiation are open and inside the target. Leaf pairs giving a very low dose remain behind the left collimator for some time at the beginning of the treatment. These leaf pairs start moving at the maximum speed once it is safe to include them in the treatment. This minimizes leaf gap effects.
■ End of delivery: All leaf pairs are at the right edge of the target.
■ At least one leaf of a pair always moves at the maximum speed to minimize the delivery time. Because of the requirements for the beginning and the end of the delivery presented above, and the differences in the complexity of the image planes, the maximum speed is individual for each leaf pair.

## Slide 59: Sliding Window

## Slide 60: Fluence Normalization
The fluence normalization is such that constant 10×10 cm2 fluence with value 1.0 provides the measured dose at 10 cm depth with the specified SPD for semi-infinite water phantom for the given amount of MU.
The SPD for the normalization geometry is the same that is used in the configuration of the treatment unit.
The normalization is based on superposition, that is, a 20×20 cm2 field with a constant fluence of 1.0 leads to roughly 1.1 times the dose (+10% of the dose) at the depth of 10 cm, while a 5×5 cm2 field gives only around 0.9. The dose values are obtained by multiplying the values by NF.
The fluence values do not compensate for the CBSF effect. CBSF is only compensated for in the MU calculation.

## Slide 61: Arc Fields in Fluence Delivery
The arc fields are modeled as a sum of individual sub-fields with a static gantry angle.
Each of these sub-fields represents a range of gantry angles in arc field.
The leaf positions and dose rate within the angular range are interpolated and the average fluence is calculated. This average fluence is then assumed to be delivered from the static gantry angle for each contributing sub-field.
The dose for the whole arc field is the sum of the doses of the static gantry angle sub-fields.