## Slide 1: Electron Cutout Factors

## Slide 2: How do you measure electron cutout factors?

## Slide 3: Electron cutout factor
How do you perform an electron MU calculation?
D0’ = 1cGy/MU at dmax, 100cm SSD, 10x10cm cone
Se = cone factor and cutout factor combined
ISL Factor is last term, where g is the gap between the treatment SSD and calibration SSD, and SSDeff is the effective SSD for the given cone size

## Slide 4: Electron cutout factor
Use the same cone size for each reading
Will dmax change with cutout inserted?
  Theoretically, yes. But the change is so small that it does not need to be accounted for
What should be used for cutout measurement?
  Small cylindrical ionization chamber, parallel plate chamber, or diode
  Minimize issues due to effective measurement depth

## Slide 5: How are electron cutout factors handled at PCAM?
Lookup table for many circular, square, and rectangular cutouts
We will typically measure electron cutout factors for cutouts that:
  Have irregular shape
  Have large disagreement between Radcalc and PennMU
  Have a dimension that is smaller than the electron range being used
  Why?
    Lack of lateral scatter equilibrium

## Slide 6: Does cutout factor change when SSD is changed?
Yes
Why?
  Each cutout creates a different virtual source position and, therefore, divergence
  At extended SSD, the relative output decreases faster for smaller cutouts

## Slide 7: Converting cutout factor at SSD1  SSD2

## Slide 8: Can cutout factor be greater than 1?
Yes
Why?
  Detector reading = Direct radiation + Head scatter radiation
    Direct radiation = Primary beam + phantom scatter hitting detector
    Head scatter radiation = Scatter from jaws, MLC, electron applicator, electron cutout
  Direct contribution always increases with increasing cutout
  Head scatter contribution depends on the energy and range of scattered particles
    Electron-electron scattering is mostly at large angles (compared to photon-electron scattering which is mostly forward)
    Scattered particles reaching the detector are mostly from the inner edge of the cutout

## Slide 9: Can cutout factor be greater than 1?

## Slide 10: Can cutout factor be greater than 1?

## Slide 11: Clinical example

## Slide 12: Clinical example