## Slide 1: Linac Components

## Slide 2: General Overview
We will trace path of power through the linac from wall  patient

## Slide 3: Questions
Is linac beam continuous or pulsed?
Why?

## Slide 4: Questions 2
How long are these pulses?
What is the time in between pulses?

## Slide 5: Modulator
Large power supply
  Supplies power to all components of the LINAC
Two main subsystems to Modulator:
  Thyratron
  Pulse Forming Network

## Slide 6: Modulator - Thyratron
Thyratron – High voltage on/off switch of the modulator
  Power requirements of linac are too great for it to run power continuously
  Capacitors are charged up and periodically discharged  controlled by thyratron
  Thyratron filled with hydrogen gas with very low pressure to allow for high speed switching
  Two separate thyratrons
    De-Qing thyratron sits in front of the Pulse Forming Network and controls the charging of its circuits
    Main thyratron sits after the PFN and controls the timing of pulses leaving the circuit
Main Thyratron

## Slide 7: Question 3
If the thyratron in your linac breaks, what symptoms will the machine exhibit?
What QA is recommended after replacement of the thyratron?

## Slide 8: Modulator
Pulse Forming Network
  Receives power from the De-Qing thyratron
  Steps up voltage from 10kV to ~100kV
  Pulses are sent to the RF driver, klystron/magnetron, and electron gun

## Slide 9: Questions
Is linac beam continuous or pulsed?
Why?

## Slide 10: Answers
Is beam from LINAC continuous or pulsed?
  Pulsed
Why?
  Heat dissipation. Very high power needs in accelerating waveguide, so by using pulses, cooling requirement is manageable

## Slide 11: Questions 2
How long are these pulses?
What is the time in between pulses?

## Slide 12: Answers 2
How long are these pulses?
What is the time in between pulses?
Duty cycle of ~0.1%

## Slide 13: Question 3
If the thyratron in your linac breaks, what symptoms will the machine exhibit?
What QA is recommended after replacement of the thyratron?

## Slide 14: Answers 3
If the thyratron in your linac breaks, what symptoms will the machine exhibit?
  There will be no beam at all as the thyratron will essentially remain in the “off switch” mode
What QA is recommended after replacement of the thyratron?
  No beam properties affected, as the thyratron is only a high voltage high speed switch. The actually voltage and pulse repetition rate is controlled by other components in the Pulse Modulator.
  Therefore, after thyratron replacement, no QA is recommended

## Slide 15: Modulator
Pulse Forming Network
  Receives power from the De-Qing thyratron
  Steps up voltage from 10kV to ~100kV
  Pulses are sent to the RF driver, klystron/magnetron, and electron gun

## Slide 16: RF Driver
Generates microwave frequency waves at 3000MHz
Operates at low power (100-300 peak watts)
Frequency of the RF driver is controlled precisely to match the resonant frequency of the main waveguide, which can vary with temperature due to expansion/contraction of the tube
  Automatic Frequency Control (AFC) monitors this

## Slide 17: Answers 4
Do all LINACs require an RF Driver?
  No
LINACs with what component require one?
  Only LINACs with klystrons require an RF Driver
  LINACs with magnetrons do not require RF Drivers

## Slide 18: Klystron
Klystron is a microwave amplifier
Needs low power microwave source input to operate
  RF driver provides this microwave input
Takes 100-300 Watt input from RF driver and steps it up to ~5-10MW
Operates like a mini-linac

## Slide 19: Klystron
Its own separate electron gun emits electrons which are injected into buncher cavity
Microwave input from RF driver is also sent into the buncher cavity
Electrons are bunched together as oscillating E-field of microwave input causes either an accelerating or retarding force on the electrons depending on when they enter the cavity
Lack of voltage difference between buncher and catcher (this region is called the drift tube) encourages the bunching effect
High power microwave output produced by passing electron bunches is extracted by the catcher cavity

## Slide 20: Question 5
What is the alternative to klystron in high-power microwave generation?
What are the advantages of this device?
What are the disadvantages?

## Slide 21: Magnetron
Can be used as a source of microwave power
  Does not need an RF driver
Functions by boiling electrons off a central filament wrapped around a cathode
Electrons are then attracted to positively charged anode on outer edge of magnetron
In the presence of external magnetic field, these electrons take on a spiral trajectory
As the electrons pass the cavity openings, a resonance phenomenon is introduced
  This leads to generation of microwaves at the resonant frequency

## Slide 22: Question 5
What is the alternative to klystron in high-power microwave generation?
What are the advantages of this device?
What are the disadvantages?

## Slide 23: Answers 5
What is the alternative to klystron in high-power microwave generation?
  Magnetron
What are the advantages of this device?
  Microwave source
  Smaller
  Cheaper
What are the disadvantages?
  Lower microwave output power
  Shorter lifespan

## Slide 24: Questions 6
What is the biggest limiting factor when considering designing a linac with a magnetron?
Which linacs use magnetron? Which use klystron?

## Slide 25: Answers 6
What is the biggest limiting factor when considering designing a linac with a magnetron?
  Reduced output power limits maximum achievable E of linac
Which linacs use magnetron? Which use klystron?
  Tomotherapy, Cyberknife both use magnetron as they only deliver 6MV. Smaller size of magnetron advantageous for those designs.
  Elekta also uses magnetron
  Varian uses klystron

## Slide 26: Circulator
One-way microwave valve
Prevents any reflected microwaves from returning to the klystron and RF driver, which would potentially damage them
Diverts reflected microwaves to a water load (for heat dissipation) or back to the accelerator waveguide so they aren’t wasted

## Slide 27: Electron Gun

## Slide 28: Electron Gun
In modern linacs, typical gun design is called a triode electron gun
  Consists of cathode, anode, and electrified grid

## Slide 29: Questions 7
What material is the filament of the electron gun?
What is the process called in which the electrons are generated from the filament?

## Slide 30: Questions 7
What material is the filament of the electron gun?
  Tungsten doped with barium oxide
    Tungsten used for its high melting point (34220C)
    Barium used as it is a good donor of electrons that are easily boiled off
What is the process called in which the electrons are generated from the filament?
  Thermionic emission  filament heated until electrons are boiled off

## Slide 31: Electron Gun
Anode is usually doughnut shaped and given a positive charge to attract electrons toward it, which then pass through the hole
Electrons then enter the accelerating waveguide at 15kV

## Slide 32: Questions 8
What is the purpose of the electrified grid?

## Slide 33: Answers 8
What is the purpose of the electrified grid?
  The grid controls when electrons are allowed to be injected into the accelerating waveguide and when they are electrically deflected away
  This feature is used to control the dose rate of the linac by controlling the frequency of electron bunches being allowed to pass through

## Slide 34: Accelerator Waveguide
Electron gun receives pulses from the modulator’s PFN in phase with the klystron output so electrons are injected into the waveguide with the proper timing for optimal acceleration

## Slide 35: Questions 9
What are the two types of accelerating waveguide designs?
Which one do most modern linacs use? Why?

## Slide 36: Accelerating Waveguide
Standing Waveguide – waves are reflected back and forth in the waveguide setting up a standing wave with high amplitude and resonance of ~3000MHz
Travelling waveguide sends microwaves down the accelerating waveguide and then either terminates the waves to a water load or recirculates them back to RF input
Standing waveguide are typically shorter in length due to constructive interference and side-coupled cavities, which move the nodes of the waves off to the side

## Slide 37: Questions 9
What are the two types of accelerating waveguide designs?
  Standing and travelling
Which one do most modern LINACs use? Why?
  Standing waveguides  Shorter, more efficient accelerators, more stable

## Slide 38: Accelerating Waveguide
Waveguide Video

## Slide 39: Standing Accelerating Waveguide
Each cavity length is ~½ wavelength
Most accelerating waveguides are operating in the S-band (3000MHz frequency)
Electrons are accelerated in each cavity. It goes through each cavity every half period.

## Slide 40: Questions 10
Are all cavity lengths actually the same?
How could you design a smaller waveguide?

## Slide 41: Answers 10
Are all cavity lengths actually the same?
  No, diameter of each cavity is the same to keep the resonance properties matched. But the earlier cavities are shorter in length as electrons are being accelerated thus gaining speed and changing the amount of time it takes to traverse a cavity
How could you design a smaller waveguide?
  Use a higher frequency (i.e. X-band = 9000MHz)
  The challenge is that it is harder to keep the waveguide stable at higher frequencies

## Slide 42: Accelerating Waveguide
The RF waveguide is filled with sulfur hexafluoride (SF6), a dielectric gas which allows higher power transmission through the waveguide without breakdown (arcing)
Accelerating waveguide is held in vacuum
Accelerating waveguide is also equipped with a solenoid coil running the length of it to keep electrons focused in a pencil beam
  Electrons naturally repel each other
Beam steering coils are used to inject the beam into the bending magnet at the proper location so that they eventually strike the target or scattering foils in the proper location

## Slide 43: Bending Magnet
Two main purposes:
  Takes beam travelling parallel to a patient and redirects towards them
  Provides energy discrimination independent of the accelerating waveguide

## Slide 44: Bending Magnet
Two varieties of bending magnet: Chromatic and Achromatic
Chromatic bend magnet uses a 900 bend. Disadvantage is that electrons emerge with different orientations depending on insertion angle, position, and energy
Achromatic designs utilize either a 2700 bend (most common) or a slalom 1200 bend to refocus the beam towards the target. Energy slit also helps with energy discrimination.

## Slide 45: Questions 11 – Target
What is the target made of? Why?
What is the main interaction that produces x-rays from the target?
What is the efficiency of this x-ray production?

## Slide 46: Answers 11 – Target
What is the target made of? Why?
  Tungsten  again mostly due to its high melting point
  Not as much due to the fact that it is high-Z
What is the main interaction that produces x-rays from the target?
  Bremsstrahlung production
What is the efficiency of this x-ray production?
  ~20% efficient

## Slide 47: Questions 12 –Target
What energy do the Bremsstrahlung photons have? Max? Average?
What is the implication of having a 20% efficiency?

## Slide 48: Answers 12 –Target
What energy do the Bremsstrahlung photons have? Max? Average?
  Max energy of the incident E
  Average energy of ~1/3 of the incident electron E
What is the implication of having a 20% efficiency?
  Need a much higher electron current for photon delivery in order to maintain the same dose rate

## Slide 49: Target
Target thickness affects….
  Beam output  thicker target = higher attenuation
  Beam spectrum  through beam hardening
  Lateral beam spread
Low energy X-rays have a different target thickness than high energy X-rays to optimize the target efficiency for each beam

## Slide 50: Primary Collimator
Round diverging collimator used to initially define the useful beam
Made of high attenuating material
What effect does the round design of the primary collimator have?

## Slide 51: Primary Collimator
Round diverging collimator used to initially define the useful beam
Made of high attenuating material
What effect does the round design of the primary collimator have?
  At largest field sizes (40x40cm2), corners of the field size are clipped and have sharper than expected falloff

## Slide 52: Questions 13
What comes first, the flattening filter or monitor chamber? Why?
What is the purpose of the flattening filter? Why is it necessary?
What is the purpose of scattering foils? Why are there two?

## Slide 53: Answers 13
What comes first, the flattening filter/scattering foils or monitor chamber? Why?
  Flattening filter/scattering foils must be first. Not only does the monitor chamber measure output, but it also measures beam flatness and symmetry
What is the purpose of the flattening filter? Why is it necessary?
  Flattening filter’s purpose is to reduce the beam fluence centrally to create a flat beam profile at a specific depth (usually 10cm)
  It is necessary because the bremsstrahlung photons created from the target are pre-dominantly forward scattered, therefore the profile before the flattening filter is peaked centrally (like an FFF beam)
What is the purpose of scattering foils? Why are there two?
  Scattering foils primary function is to scatter the electron beam to a clinically useful field size
  There are often 2 layers of scattering foils, the first layer for scattering and the second layer for flattening

## Slide 54: Flattening Filter
Cone-shaped to reduce beam fluence at center
Typically made of a medium-to-low Z material (aluminum)
  Do not want high Z material as that will significantly harden the beam
Creates flat beam profile at specific depth (10cm)
Different flattening filters are needed for each energy
Downside of flattening filters is that they further reduce photon fluence

## Slide 55: Scattering Foils
Converts thin pencil beam to a spread out electron beam
Made of lead or copper
Second foil is used to improve homogeneity across large field sizes
  First foil – scatters
  Second foil – flattens

## Slide 56: Monitor Chambers
Two layers of chambers (MU1 and MU2)
  Second layer is redundancy check in case first one fails
In each layer, 4 independent chambers are placed in each quadrant
  These are used to monitor output, flatness, and symmetry of the beam
These chambers are sealed and pressurized so environmental variations have no impact on machine output
ION1 and ION2 are separate components which deliver test charge to the MU1 and MU2 chambers before clinical beam is delivered. This test charge is used to confirm their calibration
When you change output, you are changing MU1 and MU2 by telling these chambers how much collected charge is equivalent to 1 monitor unit (based on TG-51).

## Slide 57: Remaining Beam Shaping Components
Jaws (Secondary Collimators)
  Max transmission through these collimators must be <2% (defined by the IEC)
  Clinically, transmission is typically <0.5%
  Higher the jaw position, larger the penumbra
MLCs
  MLC transmission should be less than 5% (Intraleaf leakage  usually <1%)
  Interleaf leakage (between two leaves) should be approximately 1-2%
Blocks
  Designed to have transmission of no more than 5%
  Made of cerroband
Electron cones
  Extend out from linac head to collimate beam closer to patient to minimize air gap, as electrons susceptible to scattering
Wedges

## Slide 58: Energy Selection
Three possible options
  Change electron beam current (called “beam loading”)
    The more electrons we inject into the waveguide, the less energy each one will gain from microwave input. For a fixed microwave power input, increase in electrons leads to decrease in final energy
    This technique is not used frequenty
  Change the RF output power
    Reduce RF power leaving the klystron, which will leave less energy to accelerate the electrons in the accelerating waveguide or vice versa
    This can be done by changing the RF Driver input power to the klystron or changing the klystron amplification power
  Short out part of the accelerating waveguide
    Depending on the energy needed, a certain amount of the guide is shorted out so no acceleration occurs in those cavities
    Using this technique is also called applying a plunger to the waveguide

## Slide 59: Energy Selection
  Change electron beam current (called “beam loading”)
    This technique is not used frequently
  Change the RF output power
    This technique is commonly used but there is a limited range of energies available for a fixed waveguide length
  Short out part of the accelerating waveguide
    This is another common method

## Slide 60: Energy Selection – Varian Machines
Disclaimer – I have seen conflicting information on how Varian machines operate 6e, 9e, 12e, 16e. Some sources say they operate in full waveguide mode, some say they operate in shorted mode
This info is consistent:
  There are two waveguide modes: High energy and low energy
    High energy mode utilizes the entire length of the waveguide
    Low energy mode shorts out part of the waveguide with a copper pin
    For 18X/20MeV  High energy mode is used
    For 6X  Low energy mode is used
  The klystron is fixed at a constant amplification power
  The RF driver input power is varied to the klystron for smaller energy changes (i.e. 6e9e)

## Slide 61: Energy Selection – Varian Machines
One more consideration about why plunger is definitely needed for 6X:
  Beam current required for 6X more significant than any other energy (lowest Bremsstrahlung efficiency and low output due to flattening filter)
  Higher current  less stability as well, so shorter waveguide may help with this

## Slide 62: Changing Dose Rate – Varian Machines
Varian LINACs have a fixed pulse rate
  This means the rate at which the accelerating cavity is energized is fixed. The microwave pulses will always be sent in every 4 milliseconds.
Dose rate modulation is controlled by the electrified grid of the triode electron gun
  The grid determines which pulses receive electron-loading and which do not