## Slide 1: TG 218: Tolerance limits and methodologies for IMRT measurement-based verification QA
Radiation Oncology
Daniel Alexander
(Presentation adapted from Julia Pakela and Khayrullo Shoniyozov)
July 20th, 2023

## Slide 2: History of IMRT
Conceptualized in early 90s; first clinical experiences and QA/commissioning reported in 1996
IMRT treatments feature “complex” fluence modulation; provides additional degrees of freedom with which to optimize dose distribution
Delivery methods:
  Fixed gantry (step and shoot, dynamic)
  Moving gantry (VMAT, tomotherapy)
IMRT dose distributions are typically much more heterogeneous than those of 3D plans, employing complex fields that incorporate different degrees of modulation.

## Slide 3: Purpose of patient-specific IMRT verification QA
Determine discrepancies between the planned and delivered doses
Detect gross errors in radiation delivery
‘Run through’ of the clinically delivered plan to ensure the safety of the patient
IMRT QA does not check the TPS algorithm’s management of heterogeneities, segmentation errors, or patient position errors

## Slide 4: Topics of TG 218: Review the different ways to perform IMRT QA
Review of verification metrics (dose difference, DTA, γ analysis)
Review of plan measurement methods: Composite beam, perpendicular composite, and perpendicular field by field
Review absolute dose verification methods: single point, 1D/2D analysis methodologies for absolute dose difference with ion chamber or 2D detector arrays
Review literature on expected achievable agreement between measured and calculated dose for different vendor systems and implementations (see extra slides)
Recommend Methods and Tolerance/Action Limits for Implementing IMRT QA
*Everything in TG-218 assumes a proper commissioning of your IMRT program (TG-119/120)

## Slide 5: IMRT QA - What are we analyzing?
Dose difference
Very (overly) sensitive in steep dose gradients
Distance to agreement (DTA) or contour overlay
Overly sensitive in low dose gradient regions
Dose distribution = array of points with location and dose value
  Spatial resolution does not need to be isotropic
  Interpolation occurs for displays with dose wash or isodose lines
Comparing two dose distributions – a “reference’ and an “evaluated” dose.
  Do reference and evaluated dose distributions agree to within limits that are clinically relevant?
  Some techniques are invariant with respect to the selection of the reference and evaluation distribution, others are not
  Two approaches to comparing dose distributions

## Slide 6: The dose difference test
Difference at a location () between the evaluation dose and the reference dose
Requires sampling at same grid resolution (or interpolation)
Invariant with respect to evaluated and reference (just a sign change)
Works well in low dose gradient regions
Overly sensitive in high dose gradient regions
x
D

## Slide 7: Dose to agreement (DTA)
Van Dyk et al., 1993; Harms et al., 1998
For a given point in the reference distribution, find the closest location in the evaluated dose distribution with the same dose.
  i.e
    For in :
    Find , such that &
Ideal for measuring separation between steep dose gradients
Not invariant to selection of reference distribution
Evaluated distribution usually has at least the same or greater resolution and dimensionality than the reference
Oversensitive in low dose gradient regions.
x
D

## Slide 8: Composite test
Fail
DTA
Pass
DTA
Pass ​
Dose
Pass
Test
Fail
DTA
Pass ​​
Dose
Pass ​
Test
Fail
Dose
Fail Test
Pass
DTA
Fail
Dose
Pass
Test
Takes advantage of fact that DTA and dose measurement tests are complimentary in respect to their usefulness in high/low dose gradient regions
Binary output – does not give quantitative information
  By how much did we fail?

## Slide 9: The γ (gamma) test
Generalization of composite test
  Can conceptualize as adding the normalized dose difference and DTA metrics in quadrature
The displacement between two points, , in the reference and evaluated distributions calculated in a “renormalized space” (divided by the dose () and DTA (d) criteria respectively)
Distance between the reference and evaluated point
Dose difference between reference and evaluated point
Low et al. 1998, Low et al. 2003

## Slide 10: The γ (gamma) test – visual representation

## Slide 11: Question for the audience:
When reviewing the results of a gamma analysis with a 3%/2mm criteria, you find a pair of points in the reference and evaluated distributions which yielded a dose difference of 3% and a DTA of 2 mm. What is the value of ?
≤ 1
1 < ≤ 2
> 2

## Slide 12: Question for the audience:
∆d [mm]
∆D [%]
3%
2 mm
(2 mm, 3%)
When reviewing the results of a gamma analysis with a 3%/2mm criteria, you find a pair of points in the reference and evaluated distributions which yielded a dose difference of 3% and a DTA of 2 mm. What is the value of ?
≤ 1
1 < ≤ 2
> 2
Bonus: Does this mean this point fails?

## Slide 13: Example Gamma Index Values and their relation to DTA/DD
For 3%/3mm criteria:
DTA = 0 mm | DD = 1% | γ = 0.33
DTA = 1 mm | DD = 1% | γ = 0.47
DTA = 2 mm | DD = 1% | γ = 0.75
DTA = 2 mm | DD = 2% | γ = 0.94
DTA = 3 mm | DD = 3% | γ = 1.41
DTA = 3 mm | DD = 3% | γ = 1.41
DTA = 4.5 mm | DD = 0% | γ = 1.5
In the extremes (either DTA or DD is 0), gamma index gives you the fractional value of the nonzero metric relative to its criterion

## Slide 14: The gamma test under different dose gradient conditions
In low dose gradient regions, the dose difference test dominates
In high dose gradient regions, the DTA test dominates

## Slide 15: Gamma – Strengths and Weaknesses
Strengths:
  Benefits of composite test with the addition of quantitative information
  Widely accepted QA metric across clinics
  Reduces sensitivity steep dose gradient regions (*next slide)
Limitations
  Interpretation of results is not necessarily straightforward
  Does not inherently provide site-specific feedback

## Slide 16: Dose difference vs gamma
FILM
CALCULATED
Dose Difference
For steep dose gradients with a nonzero DTA tolerance, it may not be violated for large dose differences

## Slide 17: Dose difference vs gamma
FILM
CALCULATED
GAMMA
For steep dose gradients with a nonzero DTA tolerance, it may not be violated for large dose differences

## Slide 18: Calculation parameters to consider…
Normalization
Spatial resolution
Interpretation

## Slide 19
Normalization
Global - Dose as a percentage of maximum dose for one or both of the dose distributions or as function of the prescription dose. All dose point pairs are normalized using the same value for all point pairs.
Local - the dose is normalized to the planned dose at each location for each point pair
  Allows the same relative tolerances in the target and critical structure regions
  Unrealistic dose accuracy requirements in low-dose regions
Global vs local normalization is not as sensitive within the target region
  The difference between the two is smallest in high dose areas

## Slide 20
Normalization – Question for the Audience
  Let’s say the dose is normalized to the max in the reference distribution, for example 200 cGy = 100%. Let’s say we look at a point in the low dose region of the reference distribution, which measures 20 cGy (10%), and its corresponding point in the evaluated distribution, which measures 21 cGy (10.5%). This gives a dose difference of 0.5% which will likely be acceptable. What happens if we used local normalization instead?

## Slide 21
Normalization – Question for the Audience
  Let’s say the dose is normalized to the max in the reference distribution, for example 200 cGy = 100%. Let’s say we look at a point in the low dose region of the reference distribution, which measures 20 cGy (10%), and its corresponding point in the evaluated distribution, which measures 21 cGy (10.5%). This gives a dose difference of 0.5%, which will likely be acceptable. What happens if we used local normalization instead?
  If we used local normalization, then 20 cGy = 100%, and 21 cGy = 105%, and we get a 5% dose difference. Therefore, local normalization is a more stringent test.

## Slide 22: Spatial resolution
Calculation correct
High dose gradient error
Calculation value greater
than if interpolation is used
Low dose gradient error
Spatial resolution has a big impact on accuracy of comparing dose distributions
Impacts gamma significantly for steep gradient regions (large error if no interpolation is used)
Ratio of DTA criterion to grid spacing is your limit of DTA accuracy
  Good rule of thumb, keep resolution at or below 1/3 of DTA criterion
Notes: Increase accuracy and decrease speed of the gamma calculation
Good rule of thumb is that resolution of the evaluated dose should be no greater than 1/3 of the DTA criterion (max gamma error is roughly the ratio of the spacing to the DTA criterion)

## Slide 23: Interpretation
Thresholding for gamma analysis
  Doses smaller than a certain value are not included
  It focuses on greater (clinically relevant) doses
    One result is higher passing rates
Pass/fail criteria
  Helpful when displayed as a histogram
    Is it failing in the high or low dose region?
    How many points are failing?
TG-218 Recommendations regarding Gamma Analysis:
  3%/2mm criteria
  Global Normalization
  95% tolerance limit for passing rate
  90% action limit for passing rate

## Slide 24: IMRT QA Delivery Techniques

## Slide 25: Perpendicular composite
Mounted detector perpendicular to the beam direction
Advantage: All images integrated for single dose image analysis
Con: Errors from each field may cancel on summation
Con: VMAT delivery the dose rate variation vs gantry may be obscured
Report recommends against PC

## Slide 26: Perpendicular field by field (PFF)
Gantry is fixed at 0 and collimator is fixed a at nominal angle
Advantage: Samples every part of the field, fast acquisition
Advantage: Field by field analysis can catch subtle dose differences
Con: passing rate more dependent on normalization point than for composite tests
  Question for audience: why?
Con: No 3D summation so you can’t know the significance of regional errors
TG-218 Recommends this over PC
We do this at Penn – Proton PSQA with MatrixX, IMRT QA with MapCheck and Portal Dosimetry

## Slide 27: True composite (TC)
Use real plan parameters
Advantage: uses actual gantry, collimator, couch angles. Includes inaccuracies/effects of:
  Gantry, collimator, couch angles, and MLC leaf positions with gravitational effects
  Couch top attenuation
Advantage: Provides an actual dose in a 2D plane of a 3D dose
  proximity between the high dose region and organ at risk lying on a plane can be assessed
  one dose image to analyze per plan
Con: does not sample every part of each beam if using 2D detector
  Not true if using a 3D detector
TG-218 Recommends this over others, but PFF is also good
We do this at Penn with ArcCHECK
Notes: For the film measurements, a reference field reading to a known dose, 10x10 at reference depth in a phantom
IC in a uniform dose region
Diode array
If you are using a diode array, you are worried if more than 20% of the dose comes laterally at the diode array.
Non-coplanar beams are also an issue for a diode array

## Slide 28: Absolute dose verification
Single point (small averaged volume)
  Cylindrical IC with stylized QA phantoms
  Verify MU setting to absolute dose for a point location
    Helpful for rescaling
  Measurement preformed in a non-gradient region, where dose variation is within 5% of mean chamber dose to minimize volume averaging
2D methods (single field vs composite dose)
  IC arrays, diode arrays, EPIDs, and film
  Relative dose comparisons or be calibrated and used the measured absolute dose
  Film provides the highest spatial resolution
3D methods
  3D gels
  PRESSAGE dosimeters
*Referred to TG-120 for IMRT implementation and commissioning
Notes: Worry above volume averaging in the change. 5% of the mean chamber dose … for the min and max doses across the volume of the chamber

## Slide 29: Suggestions for IMRT QA for fixed gantry and rotational
Should be performed with True Composite
  provided the QA device has negligible angular dependence or the angular dependence is accounted for
PFF if the QA device is not suitable for TC
Do not use perpendicular composite (no rotation, all summed dose)
Analysis in absolute dose mode
Recommendation of a dose calibration should be done before each measurement session
Global normalization
  should be used; deemed more clinically relevant than local normalization
  normalization point should be selected in a low gradient region with a value ≥ 90% of the max dose in the plane of measurement
Local normalization
  more stringent, only suggested during commissioning
Use a dose threshold, example of 10%
Tolerance limits for PFF and TC
  Tolerance limit: Gamma pass rate should be ≥ 95% with 3%/2mm and 10% dose threshold
  Action limit: Gamma pass rate should be ≥ 90% with 3%/2mm and 10% dose threshold

## Slide 30: Suggestions for IMRT QA for fixed gantry and rotational
If a plan fails the action limit, do not deliver the plan clinically
  Is it in a region that is clinically relevant
Tighter tolerance of 2%/1mm or 1%/1mm maybe necessary for specific treatment sites or delivery machine (SRS or SBRT)
IMRT with IC and film tolerance and action limit should be ≤2% and ≤3%
Any plan not at 100%, look at the following
  Number of points failing
  Histogram
Gamma analysis on a per structure basis
Track gamma passing rates across all patients
If your passing rates don’t match what is suggested in this document, revisit TG – 119
IROC Houston IMRT phantom suggested

## Slide 31: Actions for failed or marginal QA results
When a failure occurs, investigate it using the following:
Set Up and Beam
IMRT QA software
Multileaf Collimator (MLC)
Treatment Planning System (TPS)

## Slide 32: Actions for failed or marginal QA results
Set Up and Beam
  Phantom setup
  Correct plan version in R&V system
  Correct plan generated, correct dose per fraction, correct delivery technique, and data transfer
  Beam flatness, symmetry, and output on the day of measurement
  Beam stability (with low MU)
  Accuracy, stability, and calibration of measurement device
  Detector size and interdetector spacing (especially for SRS)
  Dose value of global normalization point
IMRT QA software
  Correct reporting and handling of plan
  Recheck Dose, DTA tolerance, threshold, registration of measured and calculation dose distributions

## Slide 33: Actions for failed or marginal QA results
Multi-leaf Collimator
  Review results of periodic general IMRT PSQA
  Leaf tolerances (speed, position, acceleration/decleration)
  Tongue and groove effects (may require high-resolution detector)
  Beam profile data for both collimator and MLC defined fields
  Dosimetric leaf gap
  Intra and Inter leaf transmission
  Jaw tracking positions (to minimize leaf transmission)

## Slide 34: Actions for failed or marginal QA results
Treatment planning system
  Amount of modulation and complexity of plan
  Number of small segments
  Total MU (affects transmission dose)
  TPS modeling accuracy for small fields (OF, profiles & penumbra)
  Characterization of leaf parameters (gap, leaf ends, transmission etc.)
  Minimum MU
  Minimum segment (or beamlet) size
  Grid size (MC)
  CT number to HU conversion
  Multiple carriage beams
  Gantry-angle spacing for VMAT delivery
  All previous items checked during IMRT commissioning

## Slide 35
Gamma pass rate per vendor were evaluated
Recommendation to vendors to work on their software and improve the additional source of uncertainty
Recommendations in TG 218 can be implemented at clinical level
Conclusions

## Slide 36: EXTRA SLIDES Reported IMRT Measured vs. Calculated Methods Agreement Vendor Survey and Algorithm Testing Process Based Tolerance and Action Limits

## Slide 37: Reported IMRT Measurement Methods
Survey of numerous institutions reported that they use:
  64% field by field and 58% absolute dose (Nelms and Simon 2007)
Criteria
  Dose difference
    76.3% of institutions used 3%
    2.9% of institutions used 4%
    15.1% of institutions used 5%
  Distance to agreement
    82.7% of institutions used 3 mm
    5.0% of institutions used 4 mm
    2.6% of institutions used 5 mm
  Threshold dose limits
    34.5% of institutions used 0-5%
    35.6% of institutions used 5-10%
    28.8% of institutions used ≥ 10%

## Slide 38
Reported IMRT Measurement
vs Calculated Agreement
IC tolerance is 5% of expected values
2D Gamma pass rates of >90% for 3%/3mm global normalization
Recent studies
  For 3%/3mm criteria, plans should have around 100% of points passing
  For 3%/2mm or 2%/2mm, plans should have pass rates with > 95% of the points passing

## Slide 39
IMRT – H&N QA
  IROC the criteria for passing is 7% dose difference reported by the TLD in the PTV and 4 mm DTA for film in the dose gradient
    Dong et al 2003- IROC for TLDS and film
      71 institutions out of 250 failed the credentialing by IROC
      71.6% passed the criteria and 28.4% failed
    Hernandez et al 2013 – update of IROC study
      156 out of 763 institutions failed credentialing by IROC
      81.6% passed the criteria and 13.7% failed only TLD, 1.8% failed the film criteria, and 2.9% failed both criteria
      This data highlights significance of IMRT verification QA
Delivery Methods - Fixed Gantry IMRT

## Slide 40
VMAT QA
  Accurate MLC motion, gantry motion, and dose rate adjustment
  Ion chamber measurements within 2.1% of MC and TPS, by Teke et al.
  MLC leaf position errors within 1 mm, 94% of the time. All errors less than 2.5 mm by Teke et al.
  Although one can calculate dose in 2 to 7 degree sectors, recommend sector calculation grid spacing of 4 degrees by Feygelmen et al
    Single position, IC measured, mean dose errors were less than 0.5% +/- 1.4% for PTV and -0.3% +/- 1.4% for OAR.
VMAT vs IMRT – no significant differences in pass rates were found when using TG-119
phantom geometry
Delivery Methods – Volumetric Modulated Arc Radiotherapy

## Slide 41
Flattening filter free
  IMRT and VMAT pass rates for Gamma of 3%/3mm are good
  Lower pass rates for a increasing ratio of MU to dose per fraction
    more MU, more highly modulated plans, thus lower pass rates
Tomotherapy
  Machine stability as reported by a study by Broggi et al
    Couch –Gantry synchronization
      Within 2mm over a 2 yr period
    Gantry MLC within 1 degree
    Dose rate variation within 1-2% during treatment
  3%/3mm gamma was found be on average 97.5%
  Correlation between IMRT QA and daily output QA for pass rates
Delivery Methods – Other modalities

## Slide 42: Considerations when using the Gamma test passing rates for evaluation
Gamma can be insensitive to finding dose errors using single field measurements and certain delivery errors
  Yan et al found that a 3%/3mm criteria and a 90% pass rate with a 2D diode array was only able to detect leaf positioning errors greater than 2 mm
  Mu et al investigated the impact of the number of control points (50 vs more than 100) on passing rates for direct machine parameter optimization (DMPO) based plans
    Worse results with more control points including D95% change between 4% and 8% and point dose to spinal cord/brainstem changed from 4% to 12%
  Kry et al looked at the number of points failed for 2%/2mm and 3%/3mm (IROC H&N)
    Fraction of pixels passing the gamma criteria from individual irradiation beams were a poor predictor of Dosimetry accuracy
  Nelms et al investigated error induced places for MLC transmission and beam penumbra
    IMRT QA was unable to show a correlation with errors and pass rates

## Slide 43: Considerations when using the Gamma test passing rates for evaluation
Take away – tighter tolerances are needed, performing a thorough analysis, and routine QA of the accelerator and MLC are required to supplement IMRT QA
Clinical analysis of IMRT QA failure results is challenging
Issues with the Gamma test
  Passing rate summarization has no spatial sensitivity
  Clustering of failed points not considered
  Large regions of low dose cause the fraction of points to appear small when the area of failed points is large compared to the high-dose regions (results in easy pass rates)

## Slide 44: Vendor Survey and Algorithm Testing

## Slide 45: Vendor Survey and Algorithm Testing

## Slide 46: Testing of vendor algorithms
Composite test and gamma index can vary from vendor to vendor
Two dose distribution provided: (A) mathematical (ideal) and (B) clinical treatment plan
Notes: In addition to the difference in QA tool designs, such as the detector type and electronics, the implementation of the same IMRT QA analysis method, such as composite test and gamma index, can vary from different vendors or even different products from the same vendors

## Slide 47: Testing of vendor algorithms
Results from different vendors [A,B,C,D,E], same plans
Ju et.al. Geometric interpretation of the gamma dose distribution comparison technique: interpolation-free calculation. Med Phys. 2008;35:879–887.

## Slide 48: Testing of vendor algorithms
Results from different vendors [A,B,C,D,E], same plans

## Slide 49: Testing of vendor algorithms
Results from different vendors [A,B,C,D,E], same plans

## Slide 50: Testing of vendor algorithms
Results from different vendors [A,B,C,D,E], same plans

## Slide 51: Testing of vendor algorithms
Mathematical test show good agreement (not perfect)
Difference in pass rates between ideal mathematical and clinical plan imply difference in dose align or handling of image resolution
Notes: In addition to the difference in QA tool designs, such as the detector type and electronics, the implementation of the same IMRT QA analysis method, such as composite test and gamma index, can vary from different vendors or even different products from the same vendors.

## Slide 52: Process Based Tolerance and Action Limits
Universally defined and guided by outcome data and expert consensus
Locally defined and guided by local experience
  Vary from institution to institution
Action Limit
    upper and lower limit of the action limit
    T is the process target value
    variance
    mean
    combination of process capability metric Cpm and error types (reject the null hypothesis if it is true).
Collect IMRT measurements over time (no out of control behavior)
  Create = moving range.
Notes: Action limit definition: Allowable deviation of measurement that won’t result in harm to the patient.
Tolerance limit definition: Are boundaries that, within, indicates that a process is subject to only random errors.

## Slide 53: Examples of Process Based Tolerance and Action Limits
VMAT and fixed gantry IMRT QA
  pass rate of 3% / 2 mm and 10% dose threshold
  evaluated with ArcCHECK and MapCHECK device
20 plans with an average gamma of 96.66% for VMAT
  Standard deviation of 1.739%
  Moving range of 1.905%
  Action limit is 22.6% according to their equation
    100-22.6%/2= 88.7%
  Control chart limit with no out of control behavior
    96.66^ - 2.660*1.905% = 91.6% (using moving range)
20 plans for with an average gamma of 95.92% for IMRT
  Standard deviation of 3.388%
  Moving range of 3.642%
  Action limit is 84.1%
  Tolerance limit is 86.2%

## Slide 54: Recommendations
Flow chart outlines the procedure for setting
tolerance and action limits for IMRT QA

## Slide 55: Passing rates for given tolerances and corresponding action limits
Notes: A number of groups have suggested metrics to assess the clinical acceptability of IMRT QA verification plans.

## Slide 56: Discussion questions
What would it look like to perform a local gamma analysis (as opposed to the usual global).
  In target region?
  In OARs?
  In high gradient regions?
Why at Penn do we choose to perform gamma analysis for portal dosimetry, but DTA for ArcCheck?
What are the benefits/limitations of using a 3D dose reconstruction algorithm for IMRT QA?